## Sheet1
PADDLE POSITION(deg) vs BAG VOLUME(l)
Test Setup
Ambu bag model: AMBU SPUR II
Ventilator->Ambu Bag->Pneumotach(Unused, but in loop)->Fluke->Lung
Bag exhaust port closed.
Run following steps for No Lung, C=50, C=20
Step 1 | Take 5 sets of readings @20 deg/sec and 100 deg/sec of the paddle
Step 2 | Average the readings
Step 3 | Curve fit the average
Step 4 | Curve fit for inverse Volume vs Degree
[16 empty] | Speed (20 deg/sec) | [6 empty] | Speed (100 deg/sec) | [6 empty] | Speed (20 deg/sec) | [6 empty] | Speed (100 deg/sec)
| Speed (20 deg/sec) | [7 empty] | Speed (100 deg/sec) | [6 empty] | With Lung, C = 50 (mL/cmH20) | [6 empty] | With Lung, C = 50 (mL/cmH20) | [6 empty] | With Lung, C = 20 (mL/cmH20) | [6 empty] | With Lung, C = 20 (mL/cmH20)
| No Lung | [7 empty] | No Lung | [7 empty] | test 1 | test 2 | test 3 | test 4 | test 5 | Average |  | test 1 | test 2 | test 3 | test 4 | test 5 | Average |  | test 1 | test 2 | test 3 | test 4 | test 5 | Average |  | test 1 | test 2 | test 3 | test 4 | test 5 | Average
|  | test 1 | test 2 | test 3 | test 4 | test 5 | Average |  |  | test 1 | test 2 | test 3 | test 4 | test 5 | Average | degrees | Volume (liters) | Volume (liters) | Volume (liters) | Volume (liters) | Volume (liters) | Volume (liters) | degrees | Volume (liters) | Volume (liters) | Volume (liters) | Volume (liters) | Volume (liters) | Volume(liters) | degrees | Volume (liters) | Volume (liters) | Volume (liters) | Volume (liters) | Volume (liters) | Volume (liters) | degrees | Volume (liters) | Volume (liters) | Volume (liters) | Volume (liters) | Volume (liters) | Volume (liters)
| degrees | Volume (liters) | Volume (liters) | Volume (liters) | Volume (liters) | Volume (liters) | Volume (liters) |  | degrees | Volume (liters) | Volume (liters) | Volume (liters) | Volume (liters) | Volume (liters) | Volume (liters) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 20 | 0.001 | 0.002 | 0.001 | 0.001 | 0.001 | 0.0012 | 20 | 0 | 0 | 0 | 0 | 0.001 | 0.0002 | 20 | 0 | 0.001 | 0 | 0.001 | 0.001 | 0.0006 | 20 | 0.001 | 0.001 | 0.001 | 0 | 0 | 0.0006
| 20 | 0.003 | 0.002 | 0.003 | 0.003 | 0.003 | 0.0028 |  | 20 | 0 | 0.001 | 0.001 | 0 | 0 | 0.0004 | 40 | 0.013 | 0.015 | 0.012 | 0.011 | 0.011 | 0.0124 | 40 | 0.004 | 0.005 | 0.004 | 0.003 | 0.004 | 0.004 | 40 | 0.003 | 0.003 | 0.002 | 0.004 | 0.003 | 0.003 | 40 | 0.004 | 0.004 | 0.003 | 0.004 | 0.004 | 0.0038
| 40 | 0.023 | 0.023 | 0.022 | 0.024 | 0.022 | 0.0228 |  | 40 | 0.011 | 0.023 | 0.021 | 0.018 | 0.019 | 0.0184 | 60 | 0.078 | 0.082 | 0.077 | 0.075 | 0.077 | 0.0778 | 60 | 0.039 | 0.047 | 0.045 | 0.042 | 0.05 | 0.0446 | 60 | 0.032 | 0.036 | 0.026 | 0.036 | 0.037 | 0.0334 | 60 | 0.038 | 0.035 | 0.035 | 0.033 | 0.036 | 0.0354
| 60 | 0.109 | 0.109 | 0.108 | 0.112 | 0.106 | 0.1088 |  | 60 | 0.055 | 0.068 | 0.07 | 0.067 | 0.069 | 0.0658 | 80 | 0.22 | 0.224 | 0.215 | 0.215 | 0.214 | 0.2176 | 80 | 0.147 | 0.156 | 0.154 | 0.151 | 0.163 | 0.1542 | 80 | 0.125 | 0.128 | 0.117 | 0.127 | 0.129 | 0.1252 | 80 | 0.128 | 0.126 | 0.0126 | 0.124 | 0.127 | 0.10352
| 80 | 0.27 | 0.273 | 0.27 | 0.275 | 0.27 | 0.2716 |  | 80 | 0.172 | 0.188 | 0.193 | 0.187 | 0.19 | 0.186 | 100 | 0.398 | 0.406 | 0.397 | 0.397 | 0.397 | 0.399 | 100 | 0.319 | 0.33 | 0.325 | 0.324 | 0.336 | 0.3268 | 100 | 0.27 | 0.271 | 0.262 | 0.271 | 0.273 | 0.2694 | 100 | 0.264 | 0.264 | 0.261 | 0.261 | 0.266 | 0.2632
| 100 | 0.469 | 0.474 | 0.47 | 0.478 | 0.47 | 0.4722 |  | 100 | 0.387 | 0.412 | 0.415 | 0.407 | 0.413 | 0.4068 | 120 | 0.611 | 0.622 | 0.612 | 0.61 | 0.611 | 0.6132 | 120 | 0.521 | 0.534 | 0.53 | 0.528 | 0.542 | 0.531 | 120 | 0.439 | 0.441 | 0.43 | 0.439 | 0.442 | 0.4382 | 120 | 0.424 | 0.424 | 0.42 | 0.422 | 0.427 | 0.4234
| 120 | 0.73 | 0.74 | 0.756 | 0.745 | 0.747 | 0.7436 |  | 120 | 0.632 | 0.657 | 0.665 | 0.661 | 0.665 | 0.656 | 140 | 0.825 | 0.842 | 0.831 | 0.824 | 0.83 | 0.8304 | 140 | 0.727 | 0.737 | 0.735 | 0.734 | 0.748 | 0.7362 | 140 | 0.607 | 0.61 | 0.597 | 0.608 | 0.61 | 0.6064 | 140 | 0.583 | 0.583 | 0.579 | 0.581 | 0.585 | 0.5822
| 140 | 1.03 | 1.028 | 1.021 | 1.027 | 1.021 | 1.0254 |  | 140 | 0.892 | 0.927 | 0.931 | 0.921 | 0.925 | 0.9192 | 160 | 0.987 | 1.004 | 0.994 | 0.989 | 0.94 | 0.9828 | 160 | 0.903 | 0.918 | 0.917 | 0.914 | 0.93 | 0.9164 | 160 | 0.738 | 0.74 | 0.726 | 0.739 | 0.74 | 0.7366 | 160 | 0.712 | 0.713 | 0.706 | 0.71 | 0.713 | 0.7108
| 160 | 1.159 | 1.144 | 1.145 | 1.146 | 1.143 | 1.1474 |  | 160 | 1.076 | 1.1 | 1.105 | 1.097 | 1.101 | 1.0958 | 180 | 1.053 | 1.071 | 1.064 | 1.055 | 1.065 | 1.0616 | 180 | 1.002 | 1.015 | 1.017 | 1.014 | 1.029 | 1.0154 | 180 | 0.815 | 0.817 | 0.801 | 0.816 | 0.813 | 0.8124 | 180 | 0.788 | 0.788 | 0.784 | 0.785 | 0.786 | 0.7862
| 180 | 1.184 | 1.178 | 1.181 | 1.183 | 1.183 | 1.1818 |  | 180 | 1.144 | 1.159 | 1.166 | 1.161 | 1.162 | 1.1584
degree
16
Volume
-6.3588e-05
1.06
[19 empty] | 1.06 | [10 empty] | 313.5134
179.2517