## 0.04 mbar steps
DIFFERENTIAL PRESSURE(mbar)vs FLOW(lpm)
Test Setup
Inspiration | Compressor->Regulator->Pneumotach(Low pressure ports of Fluke)->Fluke->Needle Valve
Expiration | Compressor->Regulator->Fluke->Pneumotach(Low pressure ports of Fluke)->Needle Valve
Change Diff Pressure(mbar) in increments of 0.04 and note Flow(lpm)
Step 1 | Take 3 sets of readings
Step 2 | Average the readings
Step 3 | Curve fit the readings
Repeat for Expiration
|  | STEP 1 - Take 3 sets of readings |  |  |  | STEP 2 - Average the readings | [12 empty] | STEP 3 - Curve fit the readings
Inflow(INSPIRATION) | [12 empty] | Average observation
Test 1 (inflow, 0.04 mbar steps) |  | Test 2 (inflow, 0.04 mbar steps) |  | Test 3(inflow, 0.04 mbar steps) |  | Average Inflow
DP (mbar) | Flow (lpm) | DP (mbar) | Flow (lpm) | DP (mbar) | Flow (lpm) | DP (mbar) | Flow (lpm) | Average Difference | Average % Difference | DP (mbar) x | Flow (lpm) y | y from curve fit | Difference | % Difference | Average Difference | Average % Difference | Result of 4 mbar step Tests
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.16473 | 2.859 | 0 | 0 | 0 | 0 | 0 | 0.16473 | 2.859 | mbar | lpm | [7 empty] | 1st Order percent error | [6 empty] | 2nd Order percent error | [6 empty] | 3rd Order percent error
0.04 | 2.7 | 0.04 | 3.08 | 0.04 | 2.9 | 0.04 | 2.8933 |  |  | 0.04 | 2.8933 | 1.7 | 1.1933 | 41.2442 |  |  | -1.92 | -88.8333 | [7 empty] | x | y from curve fit | y from fluke | Difference (fluke - curve fit) | % Change | Average Difference (lpm) | Average % Change | x | y from curve fit | y from fluke | Difference (fluke - curve fit) | % Change | Average Difference (lpm) | Average % Change | x | y from curve fit | y from fluke | Difference (fluke - curve fit) | % Change | Average Difference (lpm) | Average % Change
0.08 | 4.6 | 0.08 | 4.5 | 0.08 | 4.7 | 0.08 | 4.6 |  |  | 0.08 | 4.6 | 3.4 | 1.2 | 26.087 |  |  | -1.88 | -86.8667 | [7 empty] | -1.92 | -86.5709 | -88.8333 | -2.2625 | 2.5469 | -2.0233 | 2.4594 | -1.92 | -90.9077 | -88.8333 | 2.0744 | -2.3352 | -0.46896 | 3.3221 | -1.92 | -89.7919 | -88.8333 | 0.95853 | -1.079 | -0.42005 | 2.5158
0.12 | 6 | 0.12 | 6.14 | 0.12 | 6.2 | 0.12 | 6.1133 |  |  | 0.12 | 6.1133 | 5.1 | 1.0133 | 16.5758 |  |  | -1.84 | -85.5 | [7 empty] | -1.88 | -84.7673 | -86.8667 | -2.0993 | 2.4167 | STDEV (lpm) | STDEV | -1.88 | -88.9052 | -86.8667 | 2.0385 | -2.3467 | STDEV (lpm) | STDEV | -1.88 | -87.9358 | -86.8667 | 1.0692 | -1.2308 | STDEV (lpm) | STDEV
0.16 | 7.4 | 0.16 | 7 | 0.16 | 7.6 | 0.16 | 7.3333 |  |  | 0.16 | 7.3333 | 6.8 | 0.53333 | 7.2727 |  |  | -1.8 | -83.5 | [7 empty] | -1.84 | -82.9638 | -85.5 | -2.5362 | 2.9664 | 1.3565 | 9.5053 | -1.84 | -86.9072 | -85.5 | 1.4072 | -1.6458 | 1.036 | 8.3444 | -1.84 | -86.0765 | -85.5 | 0.57647 | -0.67424 | 1.1694 | 8.0876
0.2 | 8.7 | 0.2 | 9.1 | 0.2 | 9.1 | 0.2 | 8.9667 |  |  | 0.2 | 8.9667 | 8.5 | 0.46667 | 5.2045 |  |  | -1.76 | -82.1667 | [5 empty] | 2nd Order |  | -1.8 | -81.1602 | -83.5 | -2.3398 | 2.8022 |  |  | -1.8 | -84.9138 | -83.5 | 1.4138 | -1.6932 |  |  | -1.8 | -84.214 | -83.5 | 0.71395 | -0.85503
0.24 | 10.5 | 0.24 | 10.5 | 0.24 | 10.9 | 0.24 | 10.6333 |  |  | 0.24 | 10.6333 | 10.2 | 0.43333 | 4.0752 |  |  | -1.72 | -80.4667 |  | 1st order |  |  |  | Average Difference | % Difference | -1.76 | -79.3566 | -82.1667 | -2.81 | 3.4199 |  |  | -1.76 | -82.9251 | -82.1667 | 0.75846 | -0.92308 |  |  | -1.76 | -82.3484 | -82.1667 | 0.18178 | -0.22123
0.28 | 13.2 | 0.28 | 12.4 | 0.28 | 12.4 | 0.28 | 12.6667 |  |  | 0.28 | 12.6667 | 11.9 | 0.76667 | 6.0526 |  |  | -1.68 | -78.6 |  | Average Difference | % Difference |  |  | -0.46896 | 3.3221 | -1.72 | -77.5531 | -80.4667 | -2.9136 | 3.6209 |  |  | -1.72 | -80.941 | -80.4667 | 0.47437 | -0.58953 |  |  | -1.72 | -80.4801 | -80.4667 | 0.01347 | -0.01674
0.32 | 14.8 | 0.32 | 14.3 | 0.32 | 14.1 | 0.32 | 14.4 |  |  | 0.32 | 14.4 | 13.6 | 0.8 | 5.5556 |  |  | -1.64 | -76.4667 |  | -2.0233 | 2.4594 |  |  |  |  | -1.68 | -75.7495 | -78.6 | -2.8505 | 3.6266 |  |  | -1.68 | -78.9616 | -78.6 | 0.36157 | -0.46002 |  |  | -1.68 | -78.6092 | -78.6 | 0.0091993 | -0.011704
0.36 | 16.43 | 0.36 | 16.2 | 0.36 | 16.2 | 0.36 | 16.2767 |  |  | 0.36 | 16.2767 | 15.3 | 0.97667 | 6.0004 |  |  | -1.6 | -75.2667 | [7 empty] | -1.64 | -73.946 | -76.4667 | -2.5207 | 3.2965 |  |  | -1.64 | -76.9867 | -76.4667 | 0.52006 | -0.68012 |  |  | -1.64 | -76.7358 | -76.4667 | 0.26914 | -0.35198
0.4 | 18.2 | 0.4 | 18 | 0.4 | 17.9 | 0.4 | 18.0333 |  |  | 0.4 | 18.0333 | 17 | 1.0333 | 5.7301 |  |  | -1.56 | -73.0333 | [7 empty] | -1.6 | -72.1424 | -75.2667 | -3.1243 | 4.1509 |  |  | -1.6 | -75.0165 | -75.2667 | -0.25015 | 0.33236 |  |  | -1.6 | -74.8601 | -75.2667 | -0.40652 | 0.54011
0.44 | 19.5 | 0.44 | 19.3 | 0.44 | 19.5 | 0.44 | 19.4333 |  |  | 0.44 | 19.4333 | 18.7 | 0.73333 | 3.7736 |  |  | -1.52 | -71.2667 | [7 empty] | -1.56 | -70.3388 | -73.0333 | -2.6945 | 3.6894 |  |  | -1.56 | -73.0509 | -73.0333 | 0.017585 | -0.024079 |  |  | -1.56 | -72.9824 | -73.0333 | -0.050947 | 0.069758
0.48 | 21 | 0.48 | 21.2 | 0.48 | 21.1 | 0.48 | 21.1 |  |  | 0.48 | 21.1 | 20.4 | 0.7 | 3.3175 |  |  | -1.48 | -69.3333 | [7 empty] | -1.52 | -68.5353 | -71.2667 | -2.7314 | 3.8326 |  |  | -1.52 | -71.09 | -71.2667 | -0.17672 | 0.24797 |  |  | -1.52 | -71.1027 | -71.2667 | -0.16396 | 0.23007
0.52 | 22.8 | 0.52 | 22.7 | 0.52 | 22.9 | 0.52 | 22.8 |  |  | 0.52 | 22.8 | 22.1 | 0.7 | 3.0702 |  |  | -1.44 | -67.45 | [7 empty] | -1.48 | -66.7317 | -69.3333 | -2.6016 | 3.7523 |  |  | -1.48 | -69.1336 | -69.3333 | -0.19973 | 0.28807 |  |  | -1.48 | -69.2213 | -69.3333 | -0.11205 | 0.16161
0.56 | 24.2 | 0.56 | 24.1 | 0.56 | 24.3 | 0.56 | 24.2 |  |  | 0.56 | 24.2 | 23.8 | 0.4 | 1.6529 |  |  | -1.4 | -65.5333 | [7 empty] | -1.44 | -64.9282 | -67.45 | -2.5218 | 3.7388 |  |  | -1.44 | -67.1819 | -67.45 | -0.26811 | 0.3975 |  |  | -1.44 | -67.3383 | -67.45 | -0.11171 | 0.16562
0.6 | 25.9 | 0.6 | 25.9 | 0.6 | 25.6 | 0.6 | 25.8 |  |  | 0.6 | 25.8 | 25.5 | 0.3 | 1.1628 |  |  | -1.36 | -64.0667 | [7 empty] | -1.4 | -63.1246 | -65.5333 | -2.4087 | 3.6756 |  |  | -1.4 | -65.2348 | -65.5333 | -0.29854 | 0.45556 |  |  | -1.4 | -65.4539 | -65.5333 | -0.079425 | 0.1212
0.64 | 27.1 | 0.64 | 27.2 | 0.64 | 27 | 0.64 | 27.1 |  |  | 0.64 | 27.1 | 27.2 | -0.1 | -0.369 |  |  | -1.32 | -62 | [7 empty] | -1.36 | -61.321 | -64.0667 | -2.7456 | 4.2856 |  |  | -1.36 | -63.2923 | -64.0667 | -0.77434 | 1.2087 |  |  | -1.36 | -63.5683 | -64.0667 | -0.49835 | 0.77786
0.68 | 28.4 | 0.68 | 28.9 | 0.68 | 28.5 | 0.68 | 28.6 |  |  | 0.68 | 28.6 | 28.9 | -0.3 | -1.049 |  |  | -1.28 | -60.1367 | [7 empty] | -1.32 | -59.5175 | -62 | -2.4825 | 4.0041 |  |  | -1.32 | -61.3545 | -62 | -0.64552 | 1.0412 |  |  | -1.32 | -61.6817 | -62 | -0.31831 | 0.51341
0.72 | 29.8 | 0.72 | 30.2 | 0.72 | 30.4 | 0.72 | 30.1333 |  |  | 0.72 | 30.1333 | 30.6 | -0.46667 | -1.5487 |  |  | -1.24 | -58.4667 | [7 empty] | -1.28 | -57.7139 | -60.1367 | -2.4227 | 4.0287 |  |  | -1.28 | -59.4213 | -60.1367 | -0.71541 | 1.1896 |  |  | -1.28 | -59.7942 | -60.1367 | -0.34247 | 0.56948 |  |  | Expiration
0.76 | 31.1 | 0.76 | 31.5 | 0.76 | 31.8 | 0.76 | 31.4667 |  |  | 0.76 | 31.4667 | 32.3 | -0.83333 | -2.6483 |  |  | -1.2 | -56.6 | [7 empty] | -1.24 | -55.9104 | -58.4667 | -2.5563 | 4.3722 |  |  | -1.24 | -57.4927 | -58.4667 | -0.97401 | 1.6659 |  |  | -1.24 | -57.906 | -58.4667 | -0.56064 | 0.9589
0.8 | 32.5 | 0.8 | 33.4 | 0.8 | 32.6 | 0.8 | 32.8333 |  |  | 0.8 | 32.8333 | 34 | -1.1667 | -3.5533 |  |  | -1.16 | -54.9667 | [7 empty] | -1.2 | -54.1068 | -56.6 | -2.4932 | 4.4049 |  |  | -1.2 | -55.5687 | -56.6 | -1.0313 | 1.8221 |  |  | -1.2 | -56.0174 | -56.6 | -0.58265 | 1.0294
0.84 | 33.8 | 0.84 | 35 | 0.84 | 35.1 | 0.84 | 34.6333 |  |  | 0.84 | 34.6333 | 35.7 | -1.0667 | -3.0799 |  |  | -1.12 | -52.9 | [7 empty] | -1.16 | -52.3032 | -54.9667 | -2.6634 | 4.8455 |  |  | -1.16 | -53.6493 | -54.9667 | -1.3173 | 2.3966 |  |  | -1.16 | -54.1284 | -54.9667 | -0.83832 | 1.5251
0.88 | 35.6 | 0.88 | 36.4 | 0.88 | 36.6 | 0.88 | 36.2 |  |  | 0.88 | 36.2 | 37.4 | -1.2 | -3.3149 |  |  | -1.08 | -51.1 | [7 empty] | -1.12 | -50.4997 | -52.9 | -2.4003 | 4.5375 |  |  | -1.12 | -51.7346 | -52.9 | -1.1654 | 2.203 |  |  | -1.12 | -52.2392 | -52.9 | -0.6608 | 1.2492
0.92 | 37.3 | 0.92 | 38.4 | 0.92 | 38.7 | 0.92 | 38.1333 |  |  | 0.92 | 38.1333 | 39.1 | -0.96667 | -2.535 |  |  | -1.04 | -49.3 | [7 empty] | -1.08 | -48.6961 | -51.1 | -2.4039 | 4.7043 |  |  | -1.08 | -49.8245 | -51.1 | -1.2755 | 2.496 |  |  | -1.08 | -50.3501 | -51.1 | -0.74993 | 1.4676
0.96 | 39.6 | 0.96 | 41.1 | 0.96 | 40.4 | 0.96 | 40.3667 |  |  | 0.96 | 40.3667 | 40.8 | -0.43333 | -1.0735 |  |  | -1 | -47.2333 | [7 empty] | -1.04 | -46.8926 | -49.3 | -2.4074 | 4.8832 |  |  | -1.04 | -47.919 | -49.3 | -1.381 | 2.8011 |  |  | -1.04 | -48.4611 | -49.3 | -0.83886 | 1.7015
1 | 42 | 1 | 42.6 | 1 | 42.9 | 1 | 42.5 |  |  | 1 | 42.5 | 42.5 | 0 | 0 |  |  | -0.96 | -45.3 | [7 empty] | -1 | -45.089 | -47.2333 | -2.1443 | 4.5399 |  |  | -1 | -46.0182 | -47.2333 | -1.2151 | 2.5726 |  |  | -1 | -46.5726 | -47.2333 | -0.66073 | 1.3989
1.04 | 43.4 | 1.04 | 44.5 | 1.04 | 45 | 1.04 | 44.3 |  |  | 1.04 | 44.3 | 44.2 | 0.1 | 0.22573 |  |  | -0.92 | -43.5 | [7 empty] | -0.96 | -43.2854 | -45.3 | -2.0146 | 4.4472 |  |  | -0.96 | -44.122 | -45.3 | -1.178 | 2.6005 |  |  | -0.96 | -44.6846 | -45.3 | -0.61539 | 1.3585
1.08 | 46 | 1.08 | 47 | 1.08 | 46.5 | 1.08 | 46.5 |  |  | 1.08 | 46.5 | 45.9 | 0.6 | 1.2903 |  |  | -0.88 | -40.6 | [7 empty] | -0.92 | -41.4819 | -43.5 | -2.0181 | 4.6394 |  |  | -0.92 | -42.2304 | -43.5 | -1.2696 | 2.9187 |  |  | -0.92 | -42.7974 | -43.5 | -0.70264 | 1.6153
1.12 | 47.2 | 1.12 | 49 | 1.12 | 48 | 1.12 | 48.0667 |  |  | 1.12 | 48.0667 | 47.6 | 0.46667 | 0.97087 |  |  | -0.84 | -38.4 | [7 empty] | -0.88 | -39.6783 | -40.6 | -0.92168 | 2.2701 |  |  | -0.88 | -40.3434 | -40.6 | -0.2566 | 0.63201 |  |  | -0.88 | -40.911 | -40.6 | 0.31101 | -0.76604
1.16 | 49.2 | 1.16 | 50.5 | 1.16 | 50.4 | 1.16 | 50.0333 |  |  | 1.16 | 50.0333 | 49.3 | 0.73333 | 1.4657 |  |  | -0.8 | -36.4667 | [7 empty] | -0.84 | -37.8748 | -38.4 | -0.52524 | 1.3678 |  |  | -0.84 | -38.4611 | -38.4 | 0.061053 | -0.15899 |  |  | -0.84 | -39.0258 | -38.4 | 0.62576 | -1.6296
1.2 | 50.4 | 1.2 | 52.5 | 1.2 | 52 | 1.2 | 51.6333 |  |  | 1.2 | 51.6333 | 51 | 0.63333 | 1.2266 |  |  | -0.76 | -34.8 | [7 empty] | -0.8 | -36.0712 | -36.4667 | -0.39547 | 1.0845 |  |  | -0.8 | -36.5833 | -36.4667 | 0.11666 | -0.31991 |  |  | -0.8 | -37.1418 | -36.4667 | 0.6751 | -1.8513
1.24 | 52 | 1.24 | 54 | 1.24 | 53 | 1.24 | 53 |  |  | 1.24 | 53 | 52.7 | 0.3 | 0.56604 |  |  | -0.72 | -33.1333 | [7 empty] | -0.76 | -34.2676 | -34.8 | -0.53236 | 1.5298 |  |  | -0.76 | -34.7102 | -34.8 | -0.089772 | 0.25797 |  |  | -0.76 | -35.2592 | -34.8 | 0.45922 | -1.3196
1.28 | 54.2 | 1.28 | 55.2 | 1.28 | 55 | 1.28 | 54.8 |  |  | 1.28 | 54.8 | 54.4 | 0.4 | 0.72993 |  |  | -0.68 | -31.5667 | [7 empty] | -0.72 | -32.4641 | -33.1333 | -0.66925 | 2.0199 |  |  | -0.72 | -32.8418 | -33.1333 | -0.29158 | 0.88003 |  |  | -0.72 | -33.3783 | -33.1333 | 0.24495 | -0.73929
1.32 | 54.9 | 1.32 | 56.8 | 1.32 | 56.6 | 1.32 | 56.1 |  |  | 1.32 | 56.1 | 56.1 | 0 | -1.2666e-14 |  |  | -0.64 | -30.06 | [7 empty] | -0.68 | -30.6605 | -31.5667 | -0.90615 | 2.8706 |  |  | -0.68 | -30.9779 | -31.5667 | -0.58877 | 1.8652 |  |  | -0.68 | -31.4991 | -31.5667 | -0.067518 | 0.21389
1.36 | 57.3 | 1.36 | 59 | 1.36 | 58.4 | 1.36 | 58.2333 |  |  | 1.36 | 58.2333 | 57.8 | 0.43333 | 0.74413 |  |  | -0.6 | -28.3333 | [7 empty] | -0.64 | -28.857 | -30.06 | -1.203 | 4.0021 |  |  | -0.64 | -29.1187 | -30.06 | -0.94133 | 3.1315 |  |  | -0.64 | -29.622 | -30.06 | -0.43801 | 1.4571
1.4 | 58 | 1.4 | 60 | 1.4 | 60 | 1.4 | 59.3333 |  |  | 1.4 | 59.3333 | 59.5 | -0.16667 | -0.2809 |  |  | -0.56 | -26.8667 | [7 empty] | -0.6 | -27.0534 | -28.3333 | -1.2799 | 4.5174 |  |  | -0.6 | -27.2641 | -28.3333 | -1.0693 | 3.7739 |  |  | -0.6 | -27.747 | -28.3333 | -0.58636 | 2.0695
1.44 | 60.2 | 1.44 | 62.2 | 1.44 | 61.2 | 1.44 | 61.2 |  |  | 1.44 | 61.2 | 61.2 | 0 | 2.322e-14 |  |  | -0.52 | -25.1667 | [7 empty] | -0.56 | -25.2498 | -26.8667 | -1.6168 | 6.018 |  |  | -0.56 | -25.4141 | -26.8667 | -1.4526 | 5.4066 |  |  | -0.56 | -25.8743 | -26.8667 | -0.99238 | 3.6937
1.48 | 61.5 | 1.48 | 64.5 | 1.48 | 63.2 | 1.48 | 63.0667 |  |  | 1.48 | 63.0667 | 62.9 | 0.16667 | 0.26427 |  |  | -0.48 | -23.5 | [7 empty] | -0.52 | -23.4463 | -25.1667 | -1.7204 | 6.836 |  |  | -0.52 | -23.5687 | -25.1667 | -1.5979 | 6.3494 |  |  | -0.52 | -24.0041 | -25.1667 | -1.1626 | 4.6195
1.52 | 63 | 1.52 | 66 | 1.52 | 64.8 | 1.52 | 64.6 |  |  | 1.52 | 64.6 | 64.6 | 0 | 2.1998e-14 |  |  | -0.44 | -21.7667 | [7 empty] | -0.48 | -21.6427 | -23.5 | -1.8573 | 7.9033 |  |  | -0.48 | -21.728 | -23.5 | -1.772 | 7.5404 |  |  | -0.48 | -22.1366 | -23.5 | -1.3634 | 5.8017
1.56 | 64.5 | 1.56 | 68 | 1.56 | 66 | 1.56 | 66.1667 |  |  | 1.56 | 66.1667 | 66.3 | -0.13333 | -0.20151 |  |  | -0.4 | -20 | [7 empty] | -0.44 | -19.8392 | -21.7667 | -1.9275 | 8.8553 |  |  | -0.44 | -19.8919 | -21.7667 | -1.8748 | 8.613 |  |  | -0.44 | -20.2719 | -21.7667 | -1.4947 | 6.867
1.6 | 66 | 1.6 | 69.7 | 1.6 | 68 | 1.6 | 67.9 |  |  | 1.6 | 67.9 | 68 | -0.1 | -0.14728 |  |  | -0.36 | -18.1933 | [7 empty] | -0.4 | -18.0356 | -20 | -1.9644 | 9.822 |  |  | -0.4 | -18.0604 | -20 | -1.9396 | 9.6978 |  |  | -0.4 | -18.4103 | -20 | -1.5897 | 7.9483
1.64 | 67.4 | 1.64 | 70.7 | 1.64 | 69.5 | 1.64 | 69.2 |  |  | 1.64 | 69.2 | 69.7 | -0.5 | -0.72254 |  |  | -0.32 | -16.8333 | [7 empty] | -0.36 | -16.232 | -18.1933 | -1.9613 | 10.7803 |  |  | -0.36 | -16.2336 | -18.1933 | -1.9598 | 10.7718 |  |  | -0.36 | -16.5519 | -18.1933 | -1.6414 | 9.022
1.68 | 69.2 | 1.68 | 72.3 | 1.68 | 71 | 1.68 | 70.8333 |  |  | 1.68 | 70.8333 | 71.4 | -0.56667 | -0.8 |  |  | -0.28 | -15.1 | [7 empty] | -0.32 | -14.4285 | -16.8333 | -2.4049 | 14.2863 |  |  | -0.32 | -14.4113 | -16.8333 | -2.422 | 14.388 |  |  | -0.32 | -14.6969 | -16.8333 | -2.1364 | 12.6916
1.72 |  | 1.72 | 75 | 1.72 | 72.6 | 1.72 | 73.8 |  |  | 1.72 | 73.8 | 73.1 | 0.7 | 0.94851 |  |  | -0.24 | -13.0667 | [7 empty] | -0.28 | -12.6249 | -15.1 | -2.4751 | 16.3913 |  |  | -0.28 | -12.5937 | -15.1 | -2.5063 | 16.5977 |  |  | -0.28 | -12.8455 | -15.1 | -2.2545 | 14.9307
1.76 |  | 1.76 | 76 | 1.76 | 74 | 1.76 | 75 |  |  | 1.76 | 75 | 74.8 | 0.2 | 0.26667 |  |  | -0.2 | -10.9667 | [7 empty] | -0.24 | -10.8214 | -13.0667 | -2.2453 | 17.1835 |  |  | -0.24 | -10.7808 | -13.0667 | -2.2859 | 17.4942 |  |  | -0.24 | -10.9978 | -13.0667 | -2.0689 | 15.8335
1.8 |  | 1.8 | 78.4 | 1.8 | 76 | 1.8 | 77.2 |  |  | 1.8 | 77.2 | 76.5 | 0.7 | 0.90674 |  |  | -0.16 | -8.7 | [7 empty] | -0.2 | -9.0178 | -10.9667 | -1.9489 | 17.7708 |  |  | -0.2 | -8.9724 | -10.9667 | -1.9943 | 18.1847 |  |  | -0.2 | -9.154 | -10.9667 | -1.8127 | 16.5292
1.84 |  | 1.84 | 79.1 | 1.84 | 77.5 | 1.84 | 78.3 |  |  | 1.84 | 78.3 | 78.2 | 0.1 | 0.12771 |  |  | -0.12 | -6.9933 | [7 empty] | -0.16 | -7.2142 | -8.7 | -1.4858 | 17.0777 |  |  | -0.16 | -7.1687 | -8.7 | -1.5313 | 17.6014 |  |  | -0.16 | -7.3143 | -8.7 | -1.3857 | 15.9279
1.88 |  | 1.88 | 80.5 | 1.88 | 79.5 | 1.88 | 80 |  |  | 1.88 | 80 | 79.9 | 0.1 | 0.125 |  |  | -0.08 | -5.1833 | [7 empty] | -0.12 | -5.4107 | -6.9933 | -1.5827 | 22.6309 |  |  | -0.12 | -5.3696 | -6.9933 | -1.6238 | 23.2187 |  |  | -0.12 | -5.4789 | -6.9933 | -1.5145 | 21.656
1.92 |  | 1.92 | 82 | 1.92 | 80.5 | 1.92 | 81.25 | [9 empty] | -0.04 | -3.13 | [7 empty] | -0.08 | -3.6071 | -5.1833 | -1.5762 | 30.4093 |  |  | -0.08 | -3.5751 | -5.1833 | -1.6082 | 31.0272 |  |  | -0.08 | -3.6479 | -5.1833 | -1.5354 | 29.6228
[17 empty] | 0 | 0 | [7 empty] | -0.04 | -1.8036 | -3.13 | -1.3264 | 42.3783 |  |  | -0.04 | -1.7852 | -3.13 | -1.3448 | 42.9638 |  |  | -0.04 | -1.8215 | -3.13 | -1.3085 | 41.8038
[10 empty] | Steps of 0.04 mbar | [6 empty] | 0.04 | 2.8933 | [7 empty] | 0 | 0 | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0
[17 empty] | 0.08 | 4.6 |  | 3rd order | [5 empty] | 0.04 | 1.8036 | 2.8933 | 1.0898 | 37.665 |  |  | 0.04 | 1.7806 | 2.8933 | 1.1127 | 38.4583 |  |  | 0.04 | 1.8166 | 2.8933 | 1.0768 | 37.2157
[17 empty] | 0.12 | 6.1133 |  | Average Difference | % Difference |  |  |  |  | 0.08 | 3.6071 | 4.6 | 0.99288 | 21.5843 |  |  | 0.08 | 3.5566 | 4.6 | 1.0434 | 22.6828 |  |  | 0.08 | 3.628 | 4.6 | 0.97204 | 21.1312
[17 empty] | 0.16 | 7.3333 |  | -0.42005 | 2.5158 |  |  |  |  | 0.12 | 5.4107 | 6.1133 | 0.70265 | 11.4938 |  |  | 0.12 | 5.3279 | 6.1133 | 0.78538 | 12.8471 |  |  | 0.12 | 5.434 | 6.1133 | 0.6793 | 11.1118
[17 empty] | 0.2 | 8.9667 | [7 empty] | 0.16 | 7.2142 | 7.3333 | 0.11909 | 1.624 |  |  | 0.16 | 7.0947 | 7.3333 | 0.23865 | 3.2543 |  |  | 0.16 | 7.2346 | 7.3333 | 0.098745 | 1.3465
Outflow(EXPIRATION) | [16 empty] | 0.24 | 10.6333 |  |  |  | lpm | mbar |  |  | 0.2 | 9.0178 | 8.9667 | -0.051133 | -0.57026 |  |  | 0.2 | 8.8568 | 8.9667 | 0.10987 | 1.2254 |  |  | 0.2 | 9.0295 | 8.9667 | -0.062791 | -0.70027
Test 4 (outflow, 4 mbar steps) |  | Test 5 (outflow, 4 mbar steps) |  | Test 6 (outflow, 4 mbar steps) |  | Average Outflow | [10 empty] | 0.28 | 12.6667 |  |  |  | 30.0673 | 0.69 |  |  | 0.24 | 10.8214 | 10.6333 | -0.18803 | -1.7683 |  |  | 0.24 | 10.6143 | 10.6333 | 0.019057 | 0.17922 |  |  | 0.24 | 10.8185 | 10.6333 | -0.18513 | -1.741
[8 empty] | Average Difference | Average % Difference | [7 empty] | 0.32 | 14.4 |  |  |  | 58.7585 | 1.38 |  |  | 0.28 | 12.6249 | 12.6667 | 0.041747 | 0.32958 |  |  | 0.28 | 12.3671 | 12.6667 | 0.29953 | 2.3647 |  |  | 0.28 | 12.6014 | 12.6667 | 0.065244 | 0.51508
[8 empty] | -0.32812 | 3.6933 | [7 empty] | 0.36 | 16.2767 | [7 empty] | 0.32 | 14.4285 | 14.4 | -0.02848 | -0.19778 |  |  | 0.32 | 14.1154 | 14.4 | 0.28463 | 1.9766 |  |  | 0.32 | 14.3782 | 14.4 | 0.021835 | 0.15163
-1.92 | -89.3 | -1.92 | -88.5 | -1.92 | -88.7 | -1.92 | -88.8333 | [9 empty] | 0.4 | 18.0333 | [7 empty] | 0.36 | 16.232 | 16.2767 | 0.044627 | 0.27418 |  |  | 0.36 | 15.859 | 16.2767 | 0.41768 | 2.5662 |  |  | 0.36 | 16.1485 | 16.2767 | 0.12816 | 0.78736
-1.88 | -87.3 | -1.88 | -87 | -1.88 | -86.3 | -1.88 | -86.8667 | [9 empty] | 0.44 | 19.4333 | [7 empty] | 0.4 | 18.0356 | 18.0333 | -0.0022667 | -0.012569 |  |  | 0.4 | 17.598 | 18.0333 | 0.43537 | 2.4142 |  |  | 0.4 | 17.9123 | 18.0333 | 0.12105 | 0.67125
-1.84 | -85.8 | -1.84 | -86 | -1.84 | -84.7 | -1.84 | -85.5 |  |  | x | y from fluke | y from curve fit | Difference | % Difference | Average Difference | Average % Difference | 0.48 | 21.1 | [7 empty] | 0.44 | 19.8392 | 19.4333 | -0.40583 | -2.0883 |  |  | 0.44 | 19.3323 | 19.4333 | 0.101 | 0.51975 |  |  | 0.44 | 19.6693 | 19.4333 | -0.23597 | -1.2143
-1.8 | -83.3 | -1.8 | -84 | -1.8 | -83.2 | -1.8 | -83.5 |  |  | -1.92 | -88.8333 | -89.9155 | 1.0822 | -1.2182 | -0.32812 | 3.6933 | 0.52 | 22.8 | [7 empty] | 0.48 | 21.6427 | 21.1 | -0.54272 | -2.5721 |  |  | 0.48 | 21.0621 | 21.1 | 0.037934 | 0.17978 |  |  | 0.48 | 21.4194 | 21.1 | -0.3194 | -1.5138
-1.76 | -82.5 | -1.76 | -82.2 | -1.76 | -81.8 | -1.76 | -82.1667 |  |  | -1.88 | -86.8667 | -88.0423 | 1.1756 | -1.3534 |  |  | 0.56 | 24.2 | [7 empty] | 0.52 | 23.4463 | 22.8 | -0.64628 | -2.8346 |  |  | 0.52 | 22.7872 | 22.8 | 0.012822 | 0.056237 |  |  | 0.52 | 23.1624 | 22.8 | -0.3624 | -1.5895
-1.72 | -81 | -1.72 | -80.4 | -1.72 | -80 | -1.72 | -80.4667 |  |  | -1.84 | -85.5 | -86.169 | 0.66904 | -0.7825 |  |  | 0.6 | 25.8 | [7 empty] | 0.56 | 25.2498 | 24.2 | -1.0498 | -4.3382 |  |  | 0.56 | 24.5077 | 24.2 | -0.30767 | -1.2713 |  |  | 0.56 | 24.8981 | 24.2 | -0.69811 | -2.8847
-1.68 | -79 | -1.68 | -78.5 | -1.68 | -78.3 | -1.68 | -78.6 |  |  | -1.8 | -83.5 | -84.2958 | 0.7958 | -0.95305 |  |  | 0.64 | 27.1 | [7 empty] | 0.6 | 27.0534 | 25.8 | -1.2534 | -4.8581 |  |  | 0.6 | 26.2235 | 25.8 | -0.42353 | -1.6416 |  |  | 0.6 | 26.6264 | 25.8 | -0.82636 | -3.203
-1.64 | -77 | -1.64 | -76.4 | -1.64 | -76 | -1.64 | -76.4667 |  |  | -1.76 | -82.1667 | -82.4226 | 0.25589 | -0.31143 |  |  | 0.68 | 28.6 | [7 empty] | 0.64 | 28.857 | 27.1 | -1.757 | -6.4832 |  |  | 0.64 | 27.9348 | 27.1 | -0.83477 | -3.0803 |  |  | 0.64 | 28.347 | 27.1 | -1.247 | -4.6014
-1.6 | -75.6 | -1.6 | -76 | -1.6 | -74.2 | -1.6 | -75.2667 |  |  | -1.72 | -80.4667 | -80.5493 | 0.082653 | -0.10272 |  |  | 0.72 | 30.1333 | [7 empty] | 0.68 | 30.6605 | 28.6 | -2.0605 | -7.2046 |  |  | 0.68 | 29.6414 | 28.6 | -1.0414 | -3.6412 |  |  | 0.68 | 30.0598 | 28.6 | -1.4598 | -5.1042
-1.56 | -74 | -1.56 | -73 | -1.56 | -72.1 | -1.56 | -73.0333 |  |  | -1.68 | -78.6 | -78.6761 | 0.07608 | -0.096794 |  |  | 0.76 | 31.4667 | [7 empty] | 0.72 | 32.4641 | 30.1333 | -2.3307 | -7.7348 |  |  | 0.72 | 31.3434 | 30.1333 | -1.21 | -4.0156 |  |  | 0.72 | 31.7646 | 30.1333 | -1.6313 | -5.4135
-1.52 | -71.7 | -1.52 | -71.2 | -1.52 | -70.9 | -1.52 | -71.2667 |  |  | -1.64 | -76.4667 | -76.8028 | 0.33617 | -0.43963 |  |  | 0.8 | 32.8333 | [7 empty] | 0.76 | 34.2676 | 31.4667 | -2.801 | -8.9014 |  |  | 0.76 | 33.0407 | 31.4667 | -1.5741 | -5.0023 |  |  | 0.76 | 33.4613 | 31.4667 | -1.9946 | -6.3388 |  |  | Inspiration
-1.48 | -69.7 | -1.48 | -69.3 | -1.48 | -69 | -1.48 | -69.3333 |  |  | -1.6 | -75.2667 | -74.9296 | -0.33707 | 0.44783 |  |  | 0.84 | 34.6333 | [7 empty] | 0.8 | 36.0712 | 32.8333 | -3.2379 | -9.8615 |  |  | 0.8 | 34.7335 | 32.8333 | -1.9001 | -5.7872 |  |  | 0.8 | 35.1496 | 32.8333 | -2.3162 | -7.0545
-1.44 | -67.7 | -1.44 | -67.15 | -1.44 | -67.5 | -1.44 | -67.45 |  |  | -1.56 | -73.0333 | -73.0564 | 0.023027 | -0.031529 |  |  | 0.88 | 36.2 | [7 empty] | 0.84 | 37.8748 | 34.6333 | -3.2414 | -9.3593 |  |  | 0.84 | 36.4216 | 34.6333 | -1.7883 | -5.1634 |  |  | 0.84 | 36.8294 | 34.6333 | -2.196 | -6.3408
-1.4 | -65.6 | -1.4 | -65.2 | -1.4 | -65.8 | -1.4 | -65.5333 |  |  | -1.52 | -71.2667 | -71.1831 | -0.083547 | 0.11723 |  |  | 0.92 | 38.1333 | [7 empty] | 0.88 | 39.6783 | 36.2 | -3.4783 | -9.6086 |  |  | 0.88 | 38.1051 | 36.2 | -1.9051 | -5.2626 |  |  | 0.88 | 38.5005 | 36.2 | -2.3005 | -6.3549
-1.36 | -64 | -1.36 | -64 | -1.36 | -64.2 | -1.36 | -64.0667 |  |  | -1.48 | -69.3333 | -69.3099 | -0.023453 | 0.033827 |  |  | 0.96 | 40.3667 | [7 empty] | 0.92 | 41.4819 | 38.1333 | -3.3485 | -8.7812 |  |  | 0.92 | 39.7839 | 38.1333 | -1.6506 | -4.3285 |  |  | 0.92 | 40.1627 | 38.1333 | -2.0294 | -5.3217
-1.32 | -62.4 | -1.32 | -61.8 | -1.32 | -61.8 | -1.32 | -62 |  |  | -1.44 | -67.45 | -67.4366 | -0.01336 | 0.019807 |  |  | 1 | 42.5 | [7 empty] | 0.96 | 43.2854 | 40.3667 | -2.9188 | -7.2307 |  |  | 0.96 | 41.4582 | 40.3667 | -1.0915 | -2.704 |  |  | 0.96 | 41.8159 | 40.3667 | -1.4492 | -3.5901
-1.28 | -60 | -1.28 | -60.11 | -1.28 | -60.3 | -1.28 | -60.1367 |  |  | -1.4 | -65.5333 | -65.5634 | 0.030067 | -0.04588 |  |  | 1.04 | 44.3 | [7 empty] | 1 | 45.089 | 42.5 | -2.589 | -6.0918 |  |  | 1 | 43.1278 | 42.5 | -0.6278 | -1.4772 |  |  | 1 | 43.4598 | 42.5 | -0.9598 | -2.2584
-1.24 | -58.6 | -1.24 | -58.3 | -1.24 | -58.5 | -1.24 | -58.4667 |  |  | -1.36 | -64.0667 | -63.6902 | -0.37651 | 0.58768 |  |  | 1.08 | 46.5 | [7 empty] | 1.04 | 46.8926 | 44.3 | -2.5926 | -5.8523 |  |  | 1.04 | 44.7928 | 44.3 | -0.49279 | -1.1124 |  |  | 1.04 | 45.0943 | 44.3 | -0.79434 | -1.7931
-1.2 | -57 | -1.2 | -56.3 | -1.2 | -56.5 | -1.2 | -56.6 |  |  | -1.32 | -62 | -61.8169 | -0.18308 | 0.29529 |  |  | 1.12 | 48.0667 | [7 empty] | 1.08 | 48.6961 | 46.5 | -2.1961 | -4.7228 |  |  | 1.08 | 46.4532 | 46.5 | 0.046841 | 0.10073 |  |  | 1.08 | 46.7193 | 46.5 | -0.2193 | -0.47161
-1.16 | -55.4 | -1.16 | -55 | -1.16 | -54.5 | -1.16 | -54.9667 |  |  | -1.28 | -60.1367 | -59.9437 | -0.19299 | 0.32091 |  |  | 1.16 | 50.0333 | [7 empty] | 1.12 | 50.4997 | 48.0667 | -2.433 | -5.0617 |  |  | 1.12 | 48.1089 | 48.0667 | -0.042234 | -0.087866 |  |  | 1.12 | 48.3345 | 48.0667 | -0.26783 | -0.55721
-1.12 | -53 | -1.12 | -52.8 | -1.12 | -52.9 | -1.12 | -52.9 |  |  | -1.24 | -58.4667 | -58.0704 | -0.39623 | 0.6777 |  |  | 1.2 | 51.6333 | [7 empty] | 1.16 | 52.3032 | 50.0333 | -2.2699 | -4.5368 |  |  | 1.16 | 49.76 | 50.0333 | 0.27331 | 0.54626 |  |  | 1.16 | 49.9398 | 50.0333 | 0.093566 | 0.18701
-1.08 | -51.3 | -1.08 | -51.4 | -1.08 | -50.6 | -1.08 | -51.1 |  |  | -1.2 | -56.6 | -56.1972 | -0.4028 | 0.71166 |  |  | 1.24 | 53 | [7 empty] | 1.2 | 54.1068 | 51.6333 | -2.4735 | -4.7904 |  |  | 1.2 | 51.4065 | 51.6333 | 0.22682 | 0.43929 |  |  | 1.2 | 51.5349 | 51.6333 | 0.098412 | 0.1906
-1.04 | -49.4 | -1.04 | -49.2 | -1.04 | -49.3 | -1.04 | -49.3 |  |  | -1.16 | -54.9667 | -54.324 | -0.64271 | 1.1693 |  |  | 1.28 | 54.8 | [7 empty] | 1.24 | 55.9104 | 53 | -2.9104 | -5.4912 |  |  | 1.24 | 53.0484 | 53 | -0.04838 | -0.091284 |  |  | 1.24 | 53.1198 | 53 | -0.11979 | -0.22601
-1 | -47.5 | -1 | -47 | -1 | -47.2 | -1 | -47.2333 |  |  | -1.12 | -52.9 | -52.4507 | -0.44928 | 0.8493 |  |  | 1.32 | 56.1 | [7 empty] | 1.28 | 57.7139 | 54.8 | -2.9139 | -5.3174 |  |  | 1.28 | 54.6856 | 54.8 | 0.11438 | 0.20871 |  |  | 1.28 | 54.6942 | 54.8 | 0.10581 | 0.19309
-0.96 | -45.4 | -0.96 | -45.2 | -0.96 | -45.3 | -0.96 | -45.3 |  |  | -1.08 | -51.1 | -50.5775 | -0.52252 | 1.0225 |  |  | 1.36 | 58.2333 | [7 empty] | 1.32 | 59.5175 | 56.1 | -3.4175 | -6.0918 |  |  | 1.32 | 56.3182 | 56.1 | -0.21824 | -0.38903 |  |  | 1.32 | 56.2579 | 56.1 | -0.15794 | -0.28154
-0.92 | -43.7 | -0.92 | -43.5 | -0.92 | -43.3 | -0.92 | -43.5 |  |  | -1.04 | -49.3 | -48.7042 | -0.59576 | 1.2084 |  |  | 1.4 | 59.3333 | [7 empty] | 1.36 | 61.321 | 58.2333 | -3.0877 | -5.3023 |  |  | 1.36 | 57.9462 | 58.2333 | 0.2871 | 0.49301 |  |  | 1.36 | 57.8109 | 58.2333 | 0.42245 | 0.72545
-0.88 | -40.9 | -0.88 | -40.3 | -0.88 | -40.6 | -0.88 | -40.6 |  |  | -1 | -47.2333 | -46.831 | -0.40233 | 0.8518 |  |  | 1.44 | 61.2 | [7 empty] | 1.4 | 63.1246 | 59.3333 | -3.7913 | -6.3898 |  |  | 1.4 | 59.5696 | 59.3333 | -0.23627 | -0.39822 |  |  | 1.4 | 59.3528 | 59.3333 | -0.019487 | -0.032844
-0.84 | -38.6 | -0.84 | -38.3 | -0.84 | -38.3 | -0.84 | -38.4 |  |  | -0.96 | -45.3 | -44.9578 | -0.34224 | 0.7555 |  |  | 1.48 | 63.0667 | [7 empty] | 1.44 | 64.9282 | 61.2 | -3.7282 | -6.0918 |  |  | 1.44 | 61.1884 | 61.2 | 0.011647 | 0.019031 |  |  | 1.44 | 60.8836 | 61.2 | 0.31641 | 0.51701
-0.8 | -36.4 | -0.8 | -36.7 | -0.8 | -36.3 | -0.8 | -36.4667 |  |  | -0.92 | -43.5 | -43.0845 | -0.41548 | 0.95513 |  |  | 1.52 | 64.6 | [7 empty] | 1.48 | 66.7317 | 63.0667 | -3.6651 | -5.8114 |  |  | 1.48 | 62.8025 | 63.0667 | 0.26419 | 0.41891 |  |  | 1.48 | 62.403 | 63.0667 | 0.66366 | 1.0523
-0.76 | -34.9 | -0.76 | -34.8 | -0.76 | -34.7 | -0.76 | -34.8 |  |  | -0.88 | -40.6 | -41.2113 | 0.61128 | -1.5056 |  |  | 1.56 | 66.1667 | [7 empty] | 1.52 | 68.5353 | 64.6 | -3.9353 | -6.0918 |  |  | 1.52 | 64.412 | 64.6 | 0.18803 | 0.29107 |  |  | 1.52 | 63.9109 | 64.6 | 0.68911 | 1.0667
-0.72 | -33.12 | -0.72 | -33.2 | -0.72 | -33.08 | -0.72 | -33.1333 |  |  | -0.84 | -38.4 | -39.338 | 0.93804 | -2.4428 |  |  | 1.6 | 67.9 | [7 empty] | 1.56 | 70.3388 | 66.1667 | -4.1722 | -6.3056 |  |  | 1.56 | 66.0168 | 66.1667 | 0.14983 | 0.22644 |  |  | 1.56 | 65.4071 | 66.1667 | 0.75959 | 1.148
-0.68 | -31.6 | -0.68 | -31.5 | -0.68 | -31.6 | -0.68 | -31.5667 |  |  | -0.8 | -36.4667 | -37.4648 | 0.99813 | -2.7371 |  |  | 1.64 | 69.2 | [7 empty] | 1.6 | 72.1424 | 67.9 | -4.2424 | -6.248 |  |  | 1.6 | 67.6171 | 67.9 | 0.28291 | 0.41666 |  |  | 1.6 | 66.8914 | 67.9 | 1.0086 | 1.4855
-0.64 | -30.1 | -0.64 | -30 | -0.64 | -30.08 | -0.64 | -30.06 |  |  | -0.76 | -34.8 | -35.5916 | 0.79156 | -2.2746 |  |  | 1.68 | 70.8333 | [7 empty] | 1.64 | 73.946 | 69.2 | -4.746 | -6.8583 |  |  | 1.64 | 69.2127 | 69.2 | -0.01271 | -0.018367 |  |  | 1.64 | 68.3636 | 69.2 | 0.83638 | 1.2086
-0.6 | -28.4 | -0.6 | -28.3 | -0.6 | -28.3 | -0.6 | -28.3333 |  |  | -0.72 | -33.1333 | -33.7183 | 0.58499 | -1.7656 |  |  | 1.72 | 73.8 | [7 empty] | 1.68 | 75.7495 | 70.8333 | -4.9162 | -6.9405 |  |  | 1.68 | 70.8037 | 70.8333 | 0.029626 | 0.041825 |  |  | 1.68 | 69.8236 | 70.8333 | 1.0097 | 1.4255
-0.56 | -27 | -0.56 | -26.8 | -0.56 | -26.8 | -0.56 | -26.8667 |  |  | -0.68 | -31.5667 | -31.8451 | 0.27841 | -0.88199 |  |  | 1.76 | 75 | [7 empty] | 1.72 | 77.5531 | 73.8 | -3.7531 | -5.0855 |  |  | 1.72 | 72.3901 | 73.8 | 1.4099 | 1.9105 |  |  | 1.72 | 71.2712 | 73.8 | 2.5288 | 3.4265
-0.52 | -25.3 | -0.52 | -25.1 | -0.52 | -25.1 | -0.52 | -25.1667 |  |  | -0.64 | -30.06 | -29.9718 | -0.08816 | 0.29328 |  |  | 1.8 | 77.2 | [7 empty] | 1.76 | 79.3566 | 75 | -4.3566 | -5.8089 |  |  | 1.76 | 73.9718 | 75 | 1.0282 | 1.3709 |  |  | 1.76 | 72.7062 | 75 | 2.2938 | 3.0584
-0.48 | -23.5 | -0.48 | -23.5 | -0.48 | -23.5 | -0.48 | -23.5 |  |  | -0.6 | -28.3333 | -28.0986 | -0.23473 | 0.82847 |  |  | 1.84 | 78.3 | [7 empty] | 1.8 | 81.1602 | 77.2 | -3.9602 | -5.1298 |  |  | 1.8 | 75.549 | 77.2 | 1.651 | 2.1387 |  |  | 1.8 | 74.1285 | 77.2 | 3.0715 | 3.9787
-0.44 | -21.9 | -0.44 | -21.9 | -0.44 | -21.5 | -0.44 | -21.7667 |  |  | -0.56 | -26.8667 | -26.2254 | -0.64131 | 2.387 |  |  | 1.88 | 80 | [7 empty] | 1.84 | 82.9638 | 78.3 | -4.6638 | -5.9563 |  |  | 1.84 | 77.1215 | 78.3 | 1.1785 | 1.5052 |  |  | 1.84 | 75.5378 | 78.3 | 2.7622 | 3.5277
-0.4 | -20 | -0.4 | -20.1 | -0.4 | -19.9 | -0.4 | -20 |  |  | -0.52 | -25.1667 | -24.3521 | -0.81455 | 3.2366 |  |  | 1.92 | 81.25 | [7 empty] | 1.88 | 84.7673 | 80 | -4.7673 | -5.9591 |  |  | 1.88 | 78.6893 | 80 | 1.3107 | 1.6383 |  |  | 1.88 | 76.934 | 80 | 3.066 | 3.8326
-0.36 | -18.18 | -0.36 | -18.2 | -0.36 | -18.2 | -0.36 | -18.1933 |  |  | -0.48 | -23.5 | -22.4789 | -1.0211 | 4.3452 | [11 empty] | 1.92 | 86.5709 | 81.25 | -5.3209 | -6.5488 |  |  | 1.92 | 80.2526 | 81.25 | 0.99743 | 1.2276 |  |  | 1.92 | 78.3168 | 81.25 | 2.9332 | 3.61
-0.32 | -16.9 | -0.32 | -16.7 | -0.32 | -16.9 | -0.32 | -16.8333 |  |  | -0.44 | -21.7667 | -20.6056 | -1.161 | 5.334
-0.28 | -15.3 | -0.28 | -15.1 | -0.28 | -14.9 | -0.28 | -15.1 |  |  | -0.4 | -20 | -18.7324 | -1.2676 | 6.338
-0.24 | -13.1 | -0.24 | -13.1 | -0.24 | -13 | -0.24 | -13.0667 |  |  | -0.36 | -18.1933 | -16.8592 | -1.3342 | 7.3333
-0.2 | -11.2 | -0.2 | -10.9 | -0.2 | -10.8 | -0.2 | -10.9667 |  |  | -0.32 | -16.8333 | -14.9859 | -1.8474 | 10.9747
-0.16 | -8.8 | -0.16 | -8.7 | -0.16 | -8.6 | -0.16 | -8.7 |  |  | -0.28 | -15.1 | -13.1127 | -1.9873 | 13.1611
-0.12 | -7 | -0.12 | -6.9 | -0.12 | -7.08 | -0.12 | -6.9933 |  |  | -0.24 | -13.0667 | -11.2394 | -1.8272 | 13.9839
-0.08 | -5.3 | -0.08 | -5.15 | -0.08 | -5.1 | -0.08 | -5.1833 |  |  | -0.2 | -10.9667 | -9.3662 | -1.6005 | 14.5939
-0.04 | -3.2 | -0.04 | -3.19 | -0.04 | -3 | -0.04 | -3.13 |  |  | -0.16 | -8.7 | -7.493 | -1.207 | 13.874
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  | -0.12 | -6.9933 | -5.6197 | -1.3736 | 19.6418
[10 empty] | -0.08 | -5.1833 | -3.7465 | -1.4369 | 27.7206
[10 empty] | -0.04 | -3.13 | -1.8732 | -1.2568 | 40.1521

## NOT USED - 0.01 mbar steps
Test 1 |  |  |  | Test 3 | [5 empty] | Test 4 (back flow) |  |  |  | Test 6 (back flow)
mbar | lpm | min | max | mbar | lpm | min | max |  |  | mbar | lpm | min | max | mbar | lpm | min | max | mbar | lpm
0 | 0.445 | 0 | 0.89 | 0 | 0.385 | 0 | 0.77 |  |  | 0 | -0.58 | 0 | -1.16 | 0 | -0.43 | 0 | -0.86 | 0.3 | 13.8467
0.01 | 1.325 | 0.9 | 1.75 | 0.01 | 1.26 | 0.78 | 1.74 |  |  | -0.01 | -1.575 | -1.17 | -1.98 | -0.01 | -1.315 | -0.87 | -1.76 | 0.29 | 13.38
0.02 | 2.03 | 1.76 | 2.3 | 0.02 | 2.04 | 1.75 | 2.33 |  |  | -0.02 | -2.32 | -1.99 | -2.65 | -0.02 | -2.035 | -1.77 | -2.3 | 0.28 | 12.91
0.03 | 2.705 | 2.31 | 3.1 | 0.03 | 2.57 | 2.34 | 2.8 |  |  | -0.03 | -2.98 | -2.66 | -3.3 | -0.03 | -2.555 | -2.31 | -2.8 | 0.27 | 12.46
0.04 | 3.255 | 3.11 | 3.4 | 0.04 | 3.055 | 2.81 | 3.3 |  |  | -0.04 | -3.58 | -3.31 | -3.85 | -0.04 | -3.03 | -2.81 | -3.25 | 0.26 | 11.9133
0.05 | 3.695 | 3.41 | 3.98 | 0.05 | 3.455 | 3.31 | 3.6 |  |  | -0.05 | -4.4 | -3.86 | -4.94 | -0.05 | -3.53 | -3.26 | -3.8 | 0.25 | 11.3467
0.06 | 4.32 | 3.99 | 4.65 | 0.06 | 3.805 | 3.61 | 4 |  |  | -0.06 | -5.225 | -4.95 | -5.5 | -0.06 | -4.015 | -3.81 | -4.22 | 0.24 | 10.8433
0.07 | 4.895 | 4.66 | 5.13 | 0.07 | 4.18 | 4.01 | 4.35 |  |  | -0.07 | -5.875 | -5.51 | -6.24 | -0.07 | -4.445 | -4.23 | -4.66 | 0.23 | 10.3433
0.08 | 5.5 | 5.14 | 5.86 | 0.08 | 4.58 | 4.36 | 4.8 |  |  | -0.08 | -6.575 | -6.25 | -6.9 | -0.08 | -4.935 | -4.67 | -5.2 | 0.22 | 9.905
0.09 | 6.08 | 5.87 | 6.29 | 0.09 | 4.995 | 4.81 | 5.18 |  |  | -0.09 | -7.2 | -6.91 | -7.49 | -0.09 | -5.445 | -5.21 | -5.68 | 0.21 | 9.4783
0.1 | 6.55 | 6.3 | 6.8 | 0.1 | 5.335 | 5.19 | 5.48 |  |  | -0.1 | -7.84 | -7.5 | -8.18 | -0.1 | -5.92 | -5.69 | -6.15 | 0.2 | 9.0383
0.11 | 7.005 | 6.81 | 7.2 | 0.11 | 5.68 | 5.49 | 5.87 |  |  | -0.11 | -8.42 | -8.19 | -8.65 | -0.11 | -6.395 | -6.16 | -6.63 | 0.19 | 8.665
0.12 | 7.44 | 7.21 | 7.67 | 0.12 | 6.065 | 5.88 | 6.25 |  |  | -0.12 | -8.93 | -8.66 | -9.2 | -0.12 | -6.87 | -6.64 | -7.1 | 0.18 | 8.36
0.13 | 7.865 | 7.68 | 8.05 | 0.13 | 6.43 | 6.26 | 6.6 |  |  | -0.13 | -9.455 | -9.21 | -9.7 | -0.13 | -7.33 | -7.11 | -7.55 | 0.17 | 8.0667
0.14 | 8.32 | 8.06 | 8.58 | 0.14 | 6.755 | 6.61 | 6.9 |  |  | -0.14 | -10.105 | -9.71 | -10.5 | -0.14 | -7.78 | -7.56 | -8 | 0.16 | 7.7917
0.15 | 8.81 | 8.59 | 9.03 | 0.15 | 7.08 | 6.91 | 7.25 |  |  | -0.15 | -10.705 | -10.51 | -10.9 | -0.15 | -8.255 | -8.01 | -8.5 | 0.15 | 7.48
0.16 | 9.155 | 9.04 | 9.27 | 0.16 | 7.355 | 7.26 | 7.45 |  |  | -0.16 | -11.23 | -10.91 | -11.55 | -0.16 | -8.74 | -8.51 | -8.97 | 0.14 | 7.1067
0.17 | 9.505 | 9.28 | 9.73 | 0.17 | 7.55 | 7.46 | 7.64 |  |  | -0.17 | -11.83 | -11.56 | -12.1 | -0.17 | -9.22 | -8.98 | -9.46 | 0.13 | 6.7317
0.18 | 9.92 | 9.74 | 10.1 | 0.18 | 7.775 | 7.65 | 7.9 |  |  | -0.18 | -12.33 | -12.11 | -12.55 | -0.18 | -9.715 | -9.47 | -9.96 | 0.12 | 6.35
0.19 | 10.305 | 10.11 | 10.5 | 0.19 | 8.085 | 7.91 | 8.26 |  |  | -0.19 | -12.88 | -12.56 | -13.2 | -0.19 | -10.235 | -9.97 | -10.5 | 0.11 | 5.9717
0.2 | 10.755 | 10.51 | 11 | 0.2 | 8.51 | 8.27 | 8.75 |  |  | -0.2 | -13.48 | -13.21 | -13.75 | -0.2 | -10.79 | -10.51 | -11.07 | 0.1 | 5.595
0.21 | 11.28 | 11.01 | 11.55 | 0.21 | 8.98 | 8.76 | 9.2 |  |  | -0.21 | -14.03 | -13.76 | -14.3 | -0.21 | -11.32 | -11.08 | -11.56 | 0.09 | 5.2133
0.22 | 11.755 | 11.56 | 11.95 | 0.22 | 9.355 | 9.21 | 9.5 |  |  | -0.22 | -14.705 | -14.31 | -15.1 | -0.22 | -11.805 | -11.57 | -12.04 | 0.08 | 4.7733
0.23 | 12.195 | 11.96 | 12.43 | 0.23 | 9.755 | 9.51 | 10 |  |  | -0.23 | -15.405 | -15.11 | -15.7 | -0.23 | -12.325 | -12.05 | -12.6 | 0.07 | 4.3217
0.24 | 12.67 | 12.44 | 12.9 | 0.24 | 10.305 | 10.01 | 10.6 |  |  | -0.24 | -15.98 | -15.71 | -16.25 | -0.24 | -12.83 | -12.61 | -13.05 | 0.06 | 3.885
0.25 | 13.13 | 12.91 | 13.35 | 0.25 | 10.855 | 10.61 | 11.1 |  |  | -0.25 | -16.53 | -16.26 | -16.8 | -0.25 | -13.23 | -13.06 | -13.4 | 0.05 | 3.445
0.26 | 13.63 | 13.36 | 13.9 | 0.26 | 11.405 | 11.11 | 11.7 |  |  | -0.26 | -16.945 | -16.81 | -17.08 | -0.26 | -13.655 | -13.41 | -13.9 | 0.04 | 3.04
0.27 | 14.08 | 13.91 | 14.25 | 0.27 | 11.895 | 11.71 | 12.08 |  |  | -0.27 | -17.27 | -17.09 | -17.45 | -0.27 | -14.205 | -13.91 | -14.5 | 0.03 | 2.5517
0.28 | 14.58 | 14.26 | 14.9 | 0.28 | 12.295 | 12.09 | 12.5 |  |  | -0.28 | -17.68 | -17.46 | -17.9 | -0.28 | -14.855 | -14.51 | -15.2 | 0.02 | 1.9833
0.29 | 15.13 | 14.91 | 15.35 | 0.29 | 12.755 | 12.51 | 13 |  |  | -0.29 | -18.13 | -17.91 | -18.35 | -0.29 | -15.455 | -15.21 | -15.7 | 0.01 | 1.2617
0.3 | 15.63 | 15.36 | 15.9 | 0.3 | 13.205 | 13.01 | 13.4 |  |  | -0.3 | -18.53 | -18.36 | -18.7 | -0.3 | -15.98 | -15.71 | -16.25 | 0 | -0.059167
[18 empty] | -0.01 | -1.48
[18 empty] | -0.02 | -2.195
Test 2 |  |  |  | Average of test 1, 2, and 3 | [5 empty] | Test 5 (back flow) |  |  |  | Average of Test 4, 5, and 6 |  |  |  | -0.03 | -2.7667
mbar | lpm | min | max | mbar | lpm |  |  | Flow from Curve Fit Eq. (lpm) | Differential Pressure (mbar) | mbar | lpm | min | max | mbar | lpm |  |  | -0.04 | -3.3017
0 | 0.395 | 0 | 0.79 | 0 | 0.40833 |  |  | 5.9834 | 0.11 | 0 | -0.57 | 0 | -1.14 | 0 | -0.52667 |  |  | -0.05 | -3.9133
0.01 | 1.2 | 0.8 | 1.6 | 0.01 | 1.2617 |  |  |  |  | -0.01 | -1.55 | -1.15 | -1.95 | -0.01 | -1.48 |  |  | -0.06 | -4.5067
0.02 | 1.88 | 1.61 | 2.15 | 0.02 | 1.9833 |  |  | mbar value | psi value | -0.02 | -2.23 | -1.96 | -2.5 | -0.02 | -2.195 |  |  | -0.07 | -5.0333
0.03 | 2.38 | 2.16 | 2.6 | 0.03 | 2.5517 |  |  | 0.11 | 0.004 | -0.03 | -2.765 | -2.51 | -3.02 | -0.03 | -2.7667 |  |  | -0.08 | -5.6033
0.04 | 2.81 | 2.61 | 3.01 | 0.04 | 3.04 |  |  |  |  | -0.04 | -3.295 | -3.03 | -3.56 | -0.04 | -3.3017 |  |  | -0.09 | -6.1517
0.05 | 3.185 | 3.02 | 3.35 | 0.05 | 3.445 |  |  |  |  | -0.05 | -3.81 | -3.57 | -4.05 | -0.05 | -3.9133 |  |  | -0.1 | -6.6833
0.06 | 3.53 | 3.36 | 3.7 | 0.06 | 3.885 |  |  |  |  | -0.06 | -4.28 | -4.06 | -4.5 | -0.06 | -4.5067 |  |  | -0.11 | -7.2567
0.07 | 3.89 | 3.71 | 4.07 | 0.07 | 4.3217 |  |  |  |  | -0.07 | -4.78 | -4.51 | -5.05 | -0.07 | -5.0333 |  |  | -0.12 | -7.835
0.08 | 4.24 | 4.08 | 4.4 | 0.08 | 4.7733 |  |  |  |  | -0.08 | -5.3 | -5.06 | -5.54 | -0.08 | -5.6033 |  |  | -0.13 | -8.33
0.09 | 4.565 | 4.41 | 4.72 | 0.09 | 5.2133 |  |  |  |  | -0.09 | -5.81 | -5.55 | -6.07 | -0.09 | -6.1517 |  |  | -0.14 | -8.8467
0.1 | 4.9 | 4.73 | 5.07 | 0.1 | 5.595 |  |  |  |  | -0.1 | -6.29 | -6.08 | -6.5 | -0.1 | -6.6833 |  |  | -0.15 | -9.3683
0.11 | 5.23 | 5.08 | 5.38 | 0.11 | 5.9717 |  |  |  |  | -0.11 | -6.955 | -6.51 | -7.4 | -0.11 | -7.2567 |  |  | -0.16 | -9.8717
0.12 | 5.545 | 5.39 | 5.7 | 0.12 | 6.35 |  |  |  |  | -0.12 | -7.705 | -7.41 | -8 | -0.12 | -7.835 |  |  | -0.17 | -10.4017
0.13 | 5.9 | 5.71 | 6.09 | 0.13 | 6.7317 |  |  |  |  | -0.13 | -8.205 | -8.01 | -8.4 | -0.13 | -8.33 |  |  | -0.18 | -10.9083
0.14 | 6.245 | 6.1 | 6.39 | 0.14 | 7.1067 |  |  |  |  | -0.14 | -8.655 | -8.41 | -8.9 | -0.14 | -8.8467 |  |  | -0.19 | -11.44
0.15 | 6.55 | 6.4 | 6.7 | 0.15 | 7.48 |  |  |  |  | -0.15 | -9.145 | -8.91 | -9.38 | -0.15 | -9.3683 |  |  | -0.2 | -11.9833
0.16 | 6.865 | 6.71 | 7.02 | 0.16 | 7.7917 |  |  |  |  | -0.16 | -9.645 | -9.39 | -9.9 | -0.16 | -9.8717 |  |  | -0.21 | -12.4933
0.17 | 7.145 | 7.03 | 7.26 | 0.17 | 8.0667 |  |  |  |  | -0.17 | -10.155 | -9.91 | -10.4 | -0.17 | -10.4017 |  |  | -0.22 | -13.055
0.18 | 7.385 | 7.27 | 7.5 | 0.18 | 8.36 |  |  |  |  | -0.18 | -10.68 | -10.41 | -10.95 | -0.18 | -10.9083 |  |  | -0.23 | -13.6417
0.19 | 7.605 | 7.51 | 7.7 | 0.19 | 8.665 |  |  |  |  | -0.19 | -11.205 | -10.96 | -11.45 | -0.19 | -11.44 |  |  | -0.24 | -14.14
0.2 | 7.85 | 7.71 | 7.99 | 0.2 | 9.0383 |  |  |  |  | -0.2 | -11.68 | -11.46 | -11.9 | -0.2 | -11.9833 |  |  | -0.25 | -14.615
0.21 | 8.175 | 8 | 8.35 | 0.21 | 9.4783 |  |  |  |  | -0.21 | -12.13 | -11.91 | -12.35 | -0.21 | -12.4933 |  |  | -0.26 | -15.115
0.22 | 8.605 | 8.36 | 8.85 | 0.22 | 9.905 |  |  |  |  | -0.22 | -12.655 | -12.36 | -12.95 | -0.22 | -13.055 |  |  | -0.27 | -15.6233
0.23 | 9.08 | 8.86 | 9.3 | 0.23 | 10.3433 |  |  |  |  | -0.23 | -13.195 | -12.96 | -13.43 | -0.23 | -13.6417 |  |  | -0.28 | -16.16
0.24 | 9.555 | 9.31 | 9.8 | 0.24 | 10.8433 |  |  |  |  | -0.24 | -13.61 | -13.44 | -13.78 | -0.24 | -14.14 |  |  | -0.29 | -16.6633
0.25 | 10.055 | 9.81 | 10.3 | 0.25 | 11.3467 |  |  |  |  | -0.25 | -14.085 | -13.79 | -14.38 | -0.25 | -14.615 |  |  | -0.3 | -17.1217
0.26 | 10.705 | 10.31 | 11.1 | 0.26 | 11.9133 |  |  |  |  | -0.26 | -14.745 | -14.39 | -15.1 | -0.26 | -15.115
0.27 | 11.405 | 11.11 | 11.7 | 0.27 | 12.46 |  |  |  |  | -0.27 | -15.395 | -15.11 | -15.68 | -0.27 | -15.6233
0.28 | 11.855 | 11.71 | 12 | 0.28 | 12.91 |  |  |  |  | -0.28 | -15.945 | -15.69 | -16.2 | -0.28 | -16.16
0.29 | 12.255 | 12.01 | 12.5 | 0.29 | 13.38 |  |  |  |  | -0.29 | -16.405 | -16.21 | -16.6 | -0.29 | -16.6633
0.3 | 12.705 | 12.51 | 12.9 | 0.3 | 13.8467 |  |  |  |  | -0.3 | -16.855 | -16.61 | -17.1 | -0.3 | -17.1217